GNG2101

Design Project Progress Update

Accessible Hand Grip -AHG2

Submitted by:

Hatim Shakir, 300156065

Franck Pradel Tchombe, 300160183

Eduardo Cabada Delgado 300108419

Sebastian Schwabe, 300381061

20/10/2024

University of Ottawa

Table of Contents

Project Deliverable Report Instructions	i

Table of Contents	ii

List of Figures	iv

List of Tables	v

List of Acronyms and Glossary	vi

1	Introduction	1

2	Prototype 1, Project Progress Presentation, Peer Feedback and Team Dynamics	2

2.1	Prototype 1	2

2.2	Project Progress Presentation	2

2.3	Project plan update	2

3	Design Constraints and Prototype 2	3

3.1	Design constraints	3

3.2	Prototype 2	3

3.3	Project plan update	3

4	Economic and IP Considerations	4

4.1	Economics report	4

4.2	Intellectual property report	4

4.3	Project plan update	4

5	Design Day Pitch and Final Prototype Evaluation	5

6	Video and User Manual	6

6.1	Video pitch	6

6.2	User manual	6

7	Conclusions	7

8	Bibliography	8

List of Figures

Figure 1 Device Handler	3

Figure 2: Grip Being Help in Palm	4

Figure 3: Ergonomic shape of gripper	4

Figure 4: Gripper L Shaped Support and Two Holding Plates	5

Figure 5: Grip testing against gravity with pen	6

Figure 6: Grip Testing Against Gravity with Spoon	6

Figure 7 Roll Data Obtained from Tilting MPU from tilting MPU from 90 degrees CCW to 90 degrees CW	7

Figure 8: Pitch Data Obtained by Tilting MPU From 90 degrees CW to 90 degrees CCW	8

List of Tables

Table 1: Servo Motor Response Test to Changes in Roll and Pitch Gathered by MPU-6050	7

Table 2 Comparison table of prototype to Final Specifications	9

# Introduction

Our team is currently focused on designing an accessible hand device for the client’s daughter, who wishes to engage more independently in activities such as eating and painting. The child has difficulty holding objects like spoons or paintbrushes due to wrist positioning, so the device will aim to adjust the angle of the object they are using and improve grip for these tasks.

The first part of our work focused on identifying client needs, conducting the benchmarking process, setting target specifications, and beginning concept development. After completing phases PD B-D, we developed a detailed design and bill of materials (BOM) for our initial prototype. The next phase involves continuous prototyping and testing to ensure we meet our target specifications. Each prototype is designed to test specific functions or sub-functions while integrating our chosen design for excellence (DFX) principles.

The document is organized into five sections: prototype and project progress presentation; design constraints and prototype 2; economic and intellectual property considerations; design day pitch and final prototype; and, finally, our video demonstration and user manual. The goal of this document is to outline the various steps our team took in developing multiple prototypes and to showcase the final design that meets the target specifications, incorporates sustainability assessments, and adheres to our selected DFX principles.

# Prototype 1, Project Progress Presentation, Peer Feedback and Team Dynamics

## Prototype 1

A key challenge in our design is managing the weight and dimensions of the device. The client stressed the importance of keeping it as light as possible, so his daughter can use it effortlessly. One of our Design for Excellence (DFX) factors from Project Deliverable B was Design for Portability. Therefore, we set a goal for the device to be compact and lightweight, with a maximum length of 20 cm and a weight under 3 lbs to accommodate the mobility requirements of an 8-year-old child. After completing the benchmarking process and concept development, we established a target weight specification of less than 80 grams.

To validate this, we plan to perform weight and dimension tests. Upon building our initial prototype, we will measure the length and girth of the handle, as well as the grip diameter and the gripper’s length and width. We will use a precise scale to measure the total weight of the device and compare it against our target specification to confirm if it meets our criteria. Ensuring the device remains within these dimensions and weight limits is essential for maintaining portability and ease of use for the end user, in line with the client’s needs and our DFX objectives.

### Forearm Support

This prototype of the forearm support system will focus on the physical arm wrap that attaches to the handle that the user will grip onto. The purpose of this prototype is to test the comfort of the arm wrap and stability as the wire is very malleable and can bend when a considerable amount of weight is applied.

Figure 1. Arm wrapping prototype made from aluminum wire and soft upholstery foam made of polyurethane foam.

Figure 1 Device Handler

The photo above shows the prototype created for the forearm support arm warp, the exposed wire in the palm shows where the devices handle will be mounted.

To test this prototype for user comfort, the arm wrap will be worn for twenty minutes as this is the amount of time the client will be using the device. Comfort will be tested with the arm wrap worn directly on the arm and with a long-sleeved article of clothing. After this duration of wear, the prototype will be critiqued on any comforts or discomforts experienced. In later iterations, a fabric sleeve could be used to improve aspects of the comfort and appearance of the device as the foam will be covered and electrical wires can be routed inside the sleeve.

To test this prototype for stability, an object of known mass will be fixed to the end of the wire portion of the arm wrap to see if the arm wrap will sustain any bending caused by the weight of the rest of the device and the gripped object. The final specifications require the device to hold at most a 40-gram object with the device having a mass of at most 80 grams. Therefore, the arm wrap will be tested using a mass within the final specification, such as 120 grams, and then a mass beyond the final specification, such as 200 grams, to observe if the device is capable of more than required. If more strength is required, another length of wire can be added to the core.

The results from testing the first prototype for the arm wrap are as follows. While wearing the arm wrap, the foam provides enough padding to not feel the aluminum wire to an uncomfortable extent. When the wire is firmly wrapped around the wrist and arm, the foam provides a comfortable firm attachment to the arm with enough flexibility in the wire to not constrain movement where needed. When the arm wrap is worn and some objects of various mass are attached, the resulting effects on the wire are as follows. An object of mass 100 grams does not elastically deform the shape the wire is bent to around the arm. An object of mass 200 grams caused slight elastic deformation but does not compromise the stability of the arm wrap onto the arm.

### Grip

A critical assumption for the grip was that its ergonomic design will accommodate to the user while providing comfort during extended use. The primary purpose of this prototype was to test the grip’s ergonomics and evaluate how well it conforms to the user’s palm. It was also important to note whether it minimizes fatigue or strain during prolonged use without slipping. The prototype was 3D printed and tested by assessing comfort and grip force. Images of the grip being help in the palm, along with two picture of the design itself, demonstrate how it fits into the hand.

Figure 2: Grip Being Help in Palm

Figure 3: Ergonomic shape of gripper

### Gripper

The gripper’s design is based on the assumption that it can securely hold objects of varying cross-sectional geometries without causing slippage or damage. The prototype was developed to validate the gripper’s ability to adapt to different shapes and sizes, such as a spoon or pen, and hold them firmly using two plates held by elastic bands. The purpose of this test was to assess the grip strength and adaptability, ensuring the gripper could maintain a secure hold on objects with minimal effort. Testing involves placing various items in the gripper and analyzing how effectively they help against the force of gravity, this prototype aligns with the durability and functionality factors outlined in Deliverable B.

Figure 4: Gripper L Shaped Support and Two Holding Plates

| Figure 5: Grip testing against gravity with pen | Figure 6: Grip Testing Against Gravity with Spoon |
| --- | --- |

### Gimbal system

The prototype for this subsystem has 2 subsections, namely a hardware component and software component.

#### Hardware component

First, the electrical system in this prototype is simply the most rudimentary implementation of the basic electrical circuit that was previously designed. An Arduino Uno is synchronized with an MPU-6050 by connecting the SCL and SDA pins, and servos are connected to the Arduino Uno board through their signal pulse input wires. The servo responsible for pitch rotation is mounted on the other servo (roll rotation), which is itself housed inside the gripper. The MPU-6050 sits flush with the roll rotation servo inside the gripper as well.

A critical assumption was that this setup would fully correct against any angular displacement around y and x-axis (pitch and roll respectively). Indeed, for our device to be reliable, it needs to predictably compensate for any rotational move along those axis, which is what allows the tool to be stabilized. The purpose of this prototype was therefore to test the degrees of freedom and angular range achieved by the subsystem. It was also our goal to establish a proof-of-concept that we could fully stabilize/level an object using two servos.

To further assess the performance, the gimbal was tested in various orientations, comparing its servo response at different MPU-6050 positions. The table below shows the gimbal’s performance at the neutral horizontal positions, as well as at 90 degrees clockwise (CW) and counterclockwise (CCW) for both pitch and roll resting.

Table 1: Servo Motor Response Test to Changes in Roll and Pitch Gathered by MPU-6050

|  | Neutral Position (Horizontal) | 90 degrees CW | 90 degrees CCW |
| --- | --- | --- | --- |
| Pitch Testing |  |  |  |
| Roll Testing |  |  |  |

In Figure 6 and Figure 7, the serial plotter results are shown for pitch and roll variations to see how the servo motor arm reacts.

Figure 7 Roll Data Obtained from Tilting MPU from tilting MPU from 90 degrees CCW to 90 degrees CW

Figure 8: Pitch Data Obtained by Tilting MPU From 90 degrees CW to 90 degrees CCW

The servo motor arm covers a total range of 180 degrees in the pitch and roll direction when independently operated. This was a positive result to cover the two-DOF required.

The upcoming test will go in depth into testing different masses to test the load capacity of the gimbal system, as well as its effect on power requirements.

#### Software

The reason for splitting the gimbal system into hardware and software subparts is because of the unique set of challenges pertaining to the Arduino code itself, independently of the hardware, and because of how it directly determines the functionality of the gimbal subsystem. In this prototype, the code contains a transmission initialization section to initialize communication between the Arduino and the MPU-6050 module, a calibration section where the module is calibrated, an angle derivation (Euler angles) and filtering process (Kalman Filter) section where raw MPU-6050 data is converted into clean and relevant pitch and roll angles, and finally a servo control section, where the pitch and roll angels are fed. The purpose of this prototype was to test the angle range supported by the Euler angles model as well as the stability or smoothing provided by the Kalman Filter.

## Comparison to Final Specifications

Table 2 Comparison table of prototype to Final Specifications

| Metric # | Metric | Unit | Target Value | Test value | Test value Description |
| --- | --- | --- | --- | --- | --- |
| 1 | Total Weight | grams | <80 | <120 | Current prototype weighs less than 100 g |
| 2 | Time of Usage | Minutes | 20 | Not tested | Prototype’s battery performance has not yet been tested to determine if it can sustain use for 20 minutes. |
| 3 | Grip force | Newtons | >50 | >50 | Grip force exceeded the target threshold, demonstrating sufficient holding power for various objects |
| 4 | Angle Range | Degrees | 180 | 180 degrees | The gimbal achieved the target range of 180 degrees for pitch and roll, ensuring full rotational control |
| 5 | Heat Tolerance | °C | <52 | Not Tested | Not yet been tested to evaluate the prototype’s performance in environments with temperatures up to 52°C |
| 6 | Ease of cleaning | Binary | Yes | Not Tested | The material used (PLA) is water-resistant, but cleaning tests have not been performed |
| 7 | Noise | dB | <10 | Not Tested | -- |
| 8 | Maximum weight of gripped object | grams | <40 | Not Tested | -- |
| 9 | User interface size (Button size) | millimeters | 70-80 Diameter | N/A | Not included yet, But emergency alert may be added |
| 10 | Ambidextrous Compatible | binary | No | Yes | The way the device is held is not dependent on which hand is used |
| 11 | Wrist size range | millimeters | Approx 152.8 mm | Lateral side of hand is 3cm in OD, Medial side 4cm OD, total length= 6 cm | -- |
| 12 | Cost of device | Dollars | <$50 | <$50 | -- |
| 13 | Degrees of freedom | List | >1 | Two-DOF achieved with gimbal | -- |

## Project Progress Presentation

Project Presentation.pptx

## Project plan update

# Design Constraints and Prototype 2

## Design constraints

1. Identify your two most important non-functional design constraints (DFX factors) that play an important role in the development of your prototypes. Justify your reasoning.

The two most important non-functional design constraints (DFX factors) in this project are portability and safety. Both play a critical role in ensuring the device's usability and suitability for the our user with mobility and muscle coordination challenges.

-Design for Portability

This constraint is important because the device must be easy to wear, transport, and use without causing additional strain. A compact, lightweight design under 3 lbs and no more than 20 cm in length supports ease of movement, which is essential for the user’s comfort and functionality. Portability is further validated by testing for weight and user fatigue over 20-minute sessions, ensuring the design doesn’t cause discomfort during extended wear. Portability is key for the user who may need to use or store the device in various settings, making it practical and versatile.

Design for Safety

Given that the device is intended for a young child with impaired coordination, safety features are critical to prevent harm during use. The automatic release mechanism ensures the device can detach quickly in emergencies, which is vital for protecting a child who may not be able to remove it independently. Additional safety measures, such as rounded edges and non-slip grip surfaces, minimize injury risks and prevent accidental drops.

2. For each design constraint, explain in detail what changes would need to be made to

your design to satisfy the constraint (if any) or what already exists in the design to

respect the constraint.

-Design for Safety

To address the Design for Safety constraint, we incorporated a mechanism to prevent the device from inadvertently causing harm during a spasm. Based on the Profesoor feedback, we updated our code to recognize abnormal movements—specifically any excessive or rapid angle changes that may indicate a spasm. When such movements are detected, the code triggers an automatic stop function, halting device responsiveness to ensure safety. This approach allows the device to adapt dynamically to the user’s behavior, minimizing the risk of harm during unintended or uncontrollable movements. Additionally, we are reinforcing this with rounded edges and smooth surfaces, reducing the chance of injury from accidental contact.

-Design for Portability

For the Design for Portability constraint, we implemented a flexible, snake-like handle. This design allows the device to wrap comfortably around the user’s arm, combining an aluminum coil with a foam covering for support and comfort. The flexible handle not only secures the device in place but also ensures it remains lightweight and easy to maneuver, aligning with our goal to make the device under 3 lbs. This design respects portability requirements by enabling quick adjustments, easy storage, and a secure yet comfortable fit without restricting the user's movement.

3. Provide proof (e.g. analysis, simple calculations and/or simulations, research) to

demonstrate the effectiveness of your design in satisfying the constraints. Justify the

process and methods you used.

-Design for Safety

Automatic Stop Feature Testing: We validated the safety mechanism using code simulations that mimic unexpected spasms by introducing random, sudden angle changes. Through these simulations, we confirmed that the device responds instantly, halting movement in 100% of simulated spasm cases. This result shows the device's ability to detect and respond to abnormal behavior, critical for preventing harm.

-Design for Portability

Weight Calculations: We calculated the total weight of each component, including the aluminum coil and foam handle, to ensure the device remains under the 3 lbs limit. We also compared the weight of similar devices in related studies, which suggests a maximum weight threshold for comfortable prolonged use by children.

## Prototype 2

1. Summarize any new client feedback that you have received or any new testing results

and clearly state what needs to be changed or improved in your design. Update your

detailed design accordingly.

Prototype 1 test results indicated that elastic bands did not provide a strong grip on various objects placed within the gripper. As a solution, we implemented a set screw mechanism, illustrated in Figure 9:

Figure 9: New Gripper Design

In designing the set screw mechanism, careful consideration was given to the size of threads to ensure printability and durability, A larger thread size of M10X1 was selected to avoid potential issues with thread printing and ensure secure tightening.

Figure 10 show the grip strength testing of the new set screw design. In these tests, a 19 gram object was placed inside the gripper, the set screws were tightened, and the object was then pressed against a mass balance. Results indicate that the gripper can withstand a pulling force of up to 1.5 kg, showing a significant improvement in grip stability compared to the elastic band design.

Figure 10: Grip strength of Set Screw Testing

Gimbal

One single IMU wasn’t sufficient to fully describe both pitch and roll simultaneously. Indeed, as pitch approached ±90 degrees, the roll angle became undefined. While the system still suffers from gimbal lock, by using 2 IMUs placed perpendicularly to each other we can better describe the orientation of the device. We used the additional information provided by the second IMU to accurately map a region where the gimbal lock occurs, and while we can’t eliminate it entirely, we can make sure our application can still function despite it in a safe way. This is done by locking the device when it hits gimbal lock conditions (through software), and unlocking it automatically if and only if it is deemed it has reached a safe orientation. This prevention mechanism gets rid of any unpredictable, erratic movements, which are characteristic of gimbal locks.

In terms of software for this prototype, the objective was to improve the stability of the Kalman filter algorithm and reduce the noise further. Indeed, we considered Prof. Sykes feedback regarding the jittering of our first prototype and how it could represent a real problem. While the Kalman filter was effective, it wasn’t effective enough in reducing noise. To reduce the noise further, we accessed the DMP unit of the MPU6050 to retrieve pre-filtered sensor fusion data in quaternions, which we then converted into the familiar Euler angles needed for our application. Sensor fusion relies heavily on the gyroscope built-in the MPU6050, which is much more precise than the accelerometer, but drifts. It uses the accelerometer data mainly to correct the gyroscope’s drift, which it relies on for the actual data, leading to precise and driftless data.

2. Define the most critical product assumptions that you have not yet tested. Explain

which of your DFX factors from Project Deliverable B this assumption(s) relates to.

Forearm support

The most critical assumption for the forearm support is that it will keep the device in the user’s hands, as it was only tested by putting various masses on the end and observing how much it bends away from the open hand. To test this for the prototype 2, it will have the rest of the device attached and tested. This relates to the design for reliability, the device should remain stable in the user’s hand whether it is open or not. This also applies to the usability of the product, as it must remain usable if the user would lose grip on the handle attached to the forearm support.

Gripper

The most critical untested product assumptions for our gripper prototype center around the gripper’s reliability, usability, safety. First, we assume that the set screw mechanism will maintain a secure grip on various objects of differing sizes, shapes, and weights without loosening or failing under variable load conditions. This assumption directly relates to Reliability, as consistent grip strength is essential for the product’s intended applications. Testing will verify whether the set screw holds securely over repeated uses, especially under varying loads and torques, which is critical for product longevity.

Another assumption is that users will be able to easily adjust and tighten the set screws without requiring excessive force or fine motor skills. This assumption addresses Usability and Accessibility; ensuring that the gripper is adjustable by a wide range of users is key to making it accessible, especially for the use who has limited hand strength or dexterity. We also assume that the set screw mechanism is safe to use and does not pose any risk of injury, such as accidental pinching or exposure to sharp edges during operation. This assumption is related to Safety, as we need to confirm that the design of the set screws and their positioning minimize potential hazards. Verifying this will ensure that the gripper meets essential safety standards, especially for users who may need to adjust the screws frequently.

Finally, we assume that the material used for the set screw threads will withstand repeated tightening and loosening without degrading or stripping over time. This is linked to Material Strength and Manufacturability, as worn threads would compromise grip strength and reduce the product's lifespan. Testing the wear resistance of the printed threads will confirm whether the material can support repeated adjustments without significant degradation, which is crucial for overall product durability and user satisfaction.

Gimbal

In prototype 2, the gimbal’s main function is to test whether our device satisfies the safety and reliability DFX criteria. In prototype 1, we confirmed the gimbal could execute the expected range of motion, but what happens if the user moves the device in unexpected or extreme ways? We assumed that our device would self correct for any movement within the applicable range, but what happens if this movement exceeds this range? For this prototype, the critical assumption for the gimbal system is that it will always act predictably, even at high speeds and while withstanding unusual movements. Prototype 2 therefore implements a safety mechanism that stops the device from moving when it can no longer respond appropriately in order to avoid injuries. To improve reliability, Prototype 2 is also equipped with a more efficient digital filter that renders more accurate, denoised data, getting rid of most of the jittering that made Prototype 1 rather unreliable for activities such as eating.

3. Develop a second set of prototypes that will help you on your way to creating your final prototype and test the critical product assumptions along the way.

Forearm support

The second prototype forearm support will include a cover for the foam with electrical wires routed underneath to improve the aesthetics of this subsystem of the device. Additionally, the forearm support will be tested with the other subsystems attached instead of various objects of different mass. Another improvement that can be made on this second prototype is to wrap the foam that covers the aluminum wire core with more overlap such that bending the forearm support does not untuck the foam from under the layer above it. These improvements for the second prototype relate to the design for usability by determining how the device acts when it is assembled and maintains the usability and reliability. This also will improve the design for safety as the electrical wires will be covered by a fabric sleeve containing the forearm support and wiring.

4. Document your latest prototype(s) using as many sketches/diagrams/pictures as

required and explain the purpose and function of your prototype(s).

Forearm support

Figure: Section view of the Forearm Support for prototype 2

Figure: Basic Construction of the Fabric Forearm Support Sleeve

These two figures above show the construction of the forearm support prototype 2. The functionality is mostly visual but will provide information on comfort and routing the electrical wiring.

Gimbal and Hardware

The basic electrical circuitry is unchanged from prototype 1, except for an additional MPU6050, connected to 3.3V and GND, with SDA and SCL pins wired in parallel with the other MPU6050 module.

Figure : Circuit schematic Prototype 2

5. Carry out prototype testing, analyze and evaluate performance compared to the updated target specifications first developed in Project Deliverable C and document all your testing results and prototype specifications. Present your testing in an organized, tabular format that shows expected versus actual results (i.e. compare your measured prototype specifications to your target specifications by including both in a similar table to the one your developed for Project Deliverable C).

Table 2 Comparison table of prototype to Final Specifications

| Metric # | Metric | Unit | Final Value | Final Value Description | Design |
| --- | --- | --- | --- | --- | --- |
| 1 | Total Weight | grams | <80 | Same as ideal value from table above. | 120 |
| 2 | Time of Usage | Minutes | 20 | According to client need to perform daily activities like baking or eating that usually last about 20 mins | N/A |
| 3 | Grip force | Newtons | >50 | Same as ideal value above | 1500 |
| 4 | Angle Range | Degrees | 180 | Allow tool to face user and away from user | 180 |
| 5 | Heat Tolerance | °C | <52 | Value is same as ideal value from table above. | Not tested |
| 6 | Ease of cleaning | Binary | Yes | Material must be easy to hand-wash | Electronics involved make it difficult |
| 7 | Noise | dB | <10 | sound chart shows lower than 10 is unnoticeable [12] | <10 |
| 8 | Maximum weight of gripped object | grams | <40 | Based on average weight on object and what the object is holding (e.g. spoon with a bite of food) [13] | 1000 |
| 9 | User interface size (Button size) | millimeters | 70-80 Diameter | The client gave a qualitative comparison during meeting | 4 cm screw head |
| 10 | Ambidextrous Compatible | binary | No | Ambidextrous compatibility is not a critical requirement for our device, as it was not prioritized by the client in their specifications. | Yes |
| 11 | Wrist size range | millimeters | Approx 152.8 mm | Wrist size of average 8 year old [14] | 152.8 |
| 12 | Cost of device | Dollars | <$50 | Less than our budget. | 0 |
| 13 | Degrees of freedom | List | >1 | Requirement was for angle adjustment which only requires 1 DOF. Additional DOF may improve user comfort | 2 DOF |

Figure : Hand Curled in Position

When the hand is curled outward, the polyurethane’s snake-like flexibility offers a gentle yet supportive hold, while the servo motors adjust to keep the hand level. This dual mechanism maintains consistent control and alignment, even as the hand extends outward.

Figure : Hand Curled Back Position

In the curled-in position, the polyurethane structure provides flexible support, while two servo motors work together to maintain leveling. This setup ensures smooth movement and stability, allowing the hand to curl inward without losing alignment.

Figure : Wrist Supination

During wrist supination (palm-up position), the polyurethane material and servo motors combine to support the wrist’s rotational movement. The polyurethane provides adaptive support, and the servo motors fine-tune the alignment, minimizing strain while keeping the wrist level.

Figure : Wrist Bent Inwards

In the pronated (palm-down) position, the polyurethane structure follows the wrist’s rotation, while the servo motors actively adjust to sustain a level position. This coordinated system ensures that the hand remains steady and supported throughout the movement.

Link to Client Presentation

https://www.canva.com/design/DAGVk_LX-EA/U8c2ifoTrE5Q_wZXCkquSw/edit

Gimbal

We conducted testing comparing the new DMP filtered data to the previous accelerometer data when computing Euler angles relevant to our application (pitch and roll). This test consisted of simply observing the angle data of both methods on the Arduino IDE Serial plotter while the device was stationary and noting results.

Figure 11:Comparative angle data analysis between raw accelerometer data and DMP filtering-Pitch

Figure 12:Comparative angle data analysis between raw accelerometer data and DMP filtering-Roll

We noted that the DMP angle data had significantly less variance than the accelerometer data over the period observed. For stationary data (differentiated theoretical mean of 0), we observed that pitch accelerometer data span over 3 degrees, while DMP pitch data span less than 0.5 degrees, and roll accelerometer data span over 1 degree, while DMP roll data span less than 0.5 degrees. In both cases, the DMP data proved to be significantly less noisy, as we had expected.

## Project plan update

# Economic and IP Considerations

## Economics report

1. Include a list of variable/fixed, direct/indirect, and material/labour/overhead costs

associated with your business, based on the manufacturing and sale of your product.

Make sure that you distinguish between price and cost and realize that prototyping and

higher-volume manufacturing costs will probably be different.

2. Develop a 3-year income statement, which includes sales revenue and costs of units

sold for each year, gross profit, operating expenses and operating income (no need to

include interest and taxes).

Text here

3. Using a NPV analysis, determine the break-even point (i.e. number of units that mustbe

sold for your business to become profitable). Note: It is highly unlikely that your

operating income will be positive in the first year because of fixed costs. Therefore, you

must use a NPV analysis to compare costs and profits over multiple years based on

present value. Draw two cash flow diagrams of the expenses and incomes for the next

three years. Calculate the NPV value of each expense/income and determine the

differences and then the break-event point.

Text here

4. Describe and justify all assumptions that you have made in developing your economics

report. The assumptions must be factual based on a preliminary market research that

you conduct in order to determine the amount of demand in your target market, the

expected % of the market that you would own, and the unit price of your product based

on a sound pricing strategy.

Important note: we expect you to make many assumptions here. However, each

assumption should be identified and justified using information you gathered from vrious

sources. Provide references when using this information.

Text here

## Intellectual property report

1. Explore intellectual property databases (i.e. patents, industrial designs, integrated

circuit topographies, trademarks, copyrights, creative commons, or open source

software) to identify at least two intellectual properties related to your product or

business:

a. http://cipo.gc.ca

b. http://www.freepatentsonline.com/

c. http://patft.uspto.gov/

a. https://patents.google.com/ or other

Text here

2. Explain the importance of these intellectual properties and the legal constraints they

place on developing your product or business.

Text here

## Project plan update

Add a screenshot of your gantt chart.

# Design Day Pitch and Final Prototype Evaluation

Write your design day pitch and plan your prototype demo.

# Video and User Manual

## Video pitch

Add link to video.

## User manual

See separate template for the user manual. Do not write the content here.

# Conclusions

Summarize your lessons learned and your work related to your project. Discuss any outstanding issues or implications for the project.

# Bibliography

Insert your list of references here.