GNG2101

Design Project User and Product Manual

Waterproof Soft Cervical Collar

Submitted by:

Alien Collars Co. – B2.3

Aisven Devanand, 300143937

Gavin Garfinkel, 300354474

Matthew George, 300356554

Rayhan Karim, 300377828

Jonah Philip, 300373587

December 2nd, 2024

University of Ottawa

Table of Contents

Table of Contents	ii

List of Figures	iv

List of Tables	iv

List of Acronyms and Glossary	v

1	Introduction	1

2	Overview	2

2.1.1	Problem and Its Importance	2

2.1.2	Fundamental Needs of the User	2

2.1.3	Key Differentiators of the Product	2

2.1.4	Key Features:	3

2.2	Cautions & Warnings	4

3	Getting started	5

3.1	Configuration Considerations	6

3.2	User Access Considerations	7

3.3	Accessing/setting up the System	8

3.4	System Organization & Navigation	8

3.5	Exiting the System	9

4	Using the System	10

4.1	Function: Adjustable Support for Neck Stabilization	10

4.2	Function: Breathability Enhancement via Front Hole Design	10

4.3	Function: Rolling Up the Waterproof Soft Cervical Collar	11

5	Troubleshooting & Support	12

5.1	Error Conditions and Recovery Procedures	12

5.1.1	Loose Fit Around the Neck	12

5.1.2	Collar Does Not Stay Rolled Up	12

5.1.3	Tears or Cracks in the Silicone Coating	12

5.1.4	Foam Core Deformation	13

5.1.5	Velcro Detachment	13

5.2	Special Considerations	13

5.2.1	For Water Use	13

5.2.2	For Long-Term Storage	13

5.3	Maintenance for Error Prevention	14

5.4	Emergency Contact	14

6	Product Documentation	14

6.1.1	BOM (Bill of Materials)	14

6.1.2	Equipment list	15

6.1.3	Instructions	16

6.2	Testing & Validation	18

7	Conclusions and Recommendations for Future Work	22

8	Bibliography	23

APPENDICES	24

9	APPENDIX I: Design Files	24

List of Figures

Figure 1. Soft cervical collar - Final prototype weight	3

Figure 2. Soft cervical collar - 3 pieces of foam	7

Figure 3. Soft cervical collar - Back view	10

Figure 4. Soft cervical collar - Front view (red arrows indicate airflow)	11

Figure 5. Soft cervical collar - Dimensions	11

Figure 6. Prototypes developed for soft cervical collar (1: Left - 2: Center - 2.5: Right)	19

Figure 7. Compressibility of collar structures	21

Figure 8. Fabric layer conceptual drawing	22

List of Tables

Table 1. Acronyms	v

Table 2. Glossary	v

Table 3. Bill of Materials	14

Table 4. Summary of tests for Prototype 1	19

Table 5. Summary of tests for Prototype 2	19

Table 6. Summary of tests for Prototype 2.5	20

Table 7. Referenced Documents	24

List of Acronyms and Glossary

Table 1. Acronyms

| Acronym | Definition |
| --- | --- |

Table 2. Glossary

| Term | Acronym | Definition |
| --- | --- | --- |

# Introduction

Welcome to the user manual for the Waterproof Soft Cervical Collar, an innovative and supportive device designed for comfort and durability. This product features a core made of high-quality polyurethane foam, ensuring optimal neck support while maintaining a lightweight design. The exterior is coated with silicone, providing a water-resistant, easy-to-clean surface that enhances the product's longevity and usability in diverse conditions.

This manual is intended to guide users through the setup, usage, and maintenance of the collar, offering clear instructions and helpful visuals to ensure the best experience. Whether you're using the collar for medical recovery, posture support, or daily comfort, this document will help you maximize its benefits.

Additionally, the manual includes troubleshooting tips, detailed product specifications, and guidance for addressing potential issues. Safety precautions are also outlined to promote safe and effective use of the product.

Our goal is to provide you with a reliable, user-friendly solution to your neck support needs, ensuring your satisfaction and well-being. Thank you for choosing our Waterproof Soft Cervical Collar.

# Overview

### Problem and Its Importance

Individuals with damage to their vestibular system often experience challenges with balance, spatial orientation, and stability. For our client, who enjoys physically demanding activities such as aikido and swimming, maintaining proper neck support is critical to ensuring safety and enabling participation in these activities. Traditional cervical collars can be bulky, uncomfortable, or unsuitable for water-based activities, limiting their practicality and the user’s ability to engage in an active lifestyle.

Without an appropriate support device, the client may face risks such as further strain on their neck, discomfort during physical activities, or even accidents due to instability. This makes the development of a comfortable, water-resistant, and activity-compatible cervical collar a vital solution.

### Fundamental Needs of the User

- Effective Neck Support: The collar must provide adequate stabilization and comfort for the user's neck, especially during physically demanding activities.

- Water Resistance: As the client enjoys swimming, the collar needs to withstand water exposure without degrading or becoming uncomfortable.

- Comfort and Flexibility: The product must be lightweight and comfortable to wear for extended periods, allowing the client to move freely without irritation.

- Durability: The collar should endure repetitive use in various environments, including water and sports settings, without compromising performance.

- Ease of Maintenance: A simple cleaning process ensures the collar remains hygienic and functional over time.

### Key Differentiators of the Product

- Waterproof Design: The silicone-coated exterior makes the collar uniquely suited for water-based activities like swimming, setting it apart from conventional cervical collars.

- Lightweight and Comfortable: The use of polyurethane foam ensures a balance between support and user comfort, making it ideal for extended wear during both leisure and activity.

- Versatility for Active Lifestyles: Unlike standard cervical collars, this product is designed to accommodate the client’s participation in physically engaging activities such as aikido and swimming, without restricting movement or functionality.

- Enhanced Durability: The high-quality materials and construction ensure long-term performance even under demanding conditions.

- Enhanced Breathability: The cervical collar features a hole punch design at the front of the neck, to enhance breathability and comfort of the client, especially during hot days

Figure 1. Soft cervical collar - Final prototype weight

### Key Features:

- Supportive Core:

  - Made of high-quality polyurethane foam that maintains softness but also rigidity.

  - Lightweight design enhances user comfort for extended wear at 227 grams.

- Waterproof Exterior:

  - Silicone coating ensures water resistance, making it suitable for activities like swimming and wearing outside in the rain

  - Easy to clean and maintain, with machine washability

- Enhanced Comfort:

  - Lightweight and flexible, allowing freedom of movement.

  - Includes a hole-punch design at the front for breathability, especially in warm conditions, and Velcro for adjustability for different neck sizes

- Durability:

  - Built to withstand repetitive use and physical activities.

  - High-quality materials such as silicone ensure long-term usability.

- Portability:

- Features Velcro attachment on the front of the cervical collar, so that it can be rolled up and packed into small spaces such as a bag

## Cautions & Warnings

- Material Sensitivity:

  - Users allergic to silicone or polyurethane should avoid using the product or consult a medical professional before use.

- Restricted Activities:

  - Avoid using the collar in high-impact activities (such as boxing or wrestling) or situations beyond its intended purpose, as this may compromise its structural integrity or effectiveness, or worse: may provide injury or harm to the user.

- Extended Use:

  - Prolonged use without breaks could lead to skin irritation or discomfort. Recommended to take off the collar and check for irritation at least once every 12 hours

- Maintenance:

  - Clean the collar regularly to avoid bacterial buildup, particularly after exposure to water. Ensure there is no mold, dirt, or grease on the collar before use.

- Inspect for Damage:

  - Regularly inspect the collar for any tears, cracks, or deformities, especially after intensive use, to ensure continued effectiveness.

# Getting started

Adjusting the Collar Size and Roll for Portability

Attach the Velcro to the small Velcro square, to roll up the collar for portability

Velcro straps located on the back for adjustability and compatibility with different neck sizes

- Locate the adjustable straps or fasteners on the collar.

- Hold the collar around your neck with the opening at the front.

- Secure the straps to achieve a snug but comfortable fit:

  - Tighten slowly, ensuring there’s no discomfort or restriction of movement.

Testing Fit

- Once secured, check that the collar does not:

  - Move excessively.

  - Press against the neck uncomfortably.

- Attempt normal movements to confirm stability and that the collar supports the head.

Preparing for Use

- If the collar will be used in water:

  - Rinse with clean water to ensure the surface is free of dust or debris, use soap to rinse off grease if needed, if the collar has been used for an extended period, wash in the washing machine as needed

## Configuration Considerations

- Collar is comprised of three pieces, as seen in Figure 2

  - Cut polyurethane foam into 3 pieces, the main central piece which is placed under the neck with three holes, and two side pieces that support the side and back of the neck. Measurements are subject to client preferences and measurements of the client’s neck

- Pieces connected with adhesive

  - The pieces are glued together with a strong adhesive. Two-part epoxy was used for the original prototype, however any adhesive that is strong enough will work

- Collar coated with silicone

  - The collar has been coated in a silicone layer for the waterproofing

- Velcro

  - Velcro has been attached to the cervical collar via threads and sewing threads into the Velcro to securely attach it to the collar. Velcro is featured at the back for size adjustability, and there is a single Velcro patch on the central piece which is to allow the user to roll up the cervical collar into a more space efficient shape.

Figure 2. Soft cervical collar - 3 pieces of foam

## User Access Considerations

- Individuals Recovering from Neck Injuries

  - Purpose: To provide stabilization and support during recovery from minor cervical injuries or surgeries.

  - Restrictions:

    - Must be used under medical supervision to ensure proper fit and suitability.

    - Not intended for individuals with severe neck or spinal injuries requiring rigid and very firm neck support.

- Active Individuals with Neck Support Needs

  - Purpose: To aid those engaging in physical activities like swimming, martial arts, or yoga while requiring neck support.

  - Restrictions:

    - Avoid activities involving high impact or excessive strain on the neck.

    - Users should monitor for discomfort or reduced performance due to prolonged use.

- People with Postural Issues

  - Purpose: To assist individuals in maintaining proper neck posture during daily activities or desk work.

  - Restrictions:

    - Limit the use to specific durations as over-reliance might weaken neck muscles.

    - Ensure the collar does not interfere with tasks requiring full neck mobility.

- Elderly Individuals or Those with Vestibular Issues

  - Purpose: To provide support for balance and stability, especially during walking or light exercises.

  - Restrictions:

    - Supervision may be necessary for elderly users to ensure correct placement and usage.

    - Not designed as a substitute for more comprehensive mobility aids.

## Accessing/setting up the System

- Position the Collar:

  - Hold the collar with the opening facing forward.

  - Align the center of the back portion with the nape of your neck.

- Secure the Collar:

  - Wrap the collar snugly around your neck.

  - Attach the straps or Velcro fasteners at the back to secure it in place, adjust until the collar hugs the neck to the desired firmness.

  - Avoid over-tightening; ensure that the collar is tight enough that it supports the neck and allows for neck mobility, never overtighten so that the collar restricts blood flow

- Adjust for Comfort:

  - Ensure the collar fits evenly and does not tilt or cause uncomfortable pressure points.

  - Confirm that the hole-punch design aligns with your neck for airflow.

## System Organization & Navigation

- Core Structure:

  - Material: High-quality polyurethane foam.

  - Function: Provides the foundational support for the collar, ensuring it is lightweight yet sturdy enough to stabilize the neck.

- Waterproof Layer:

  - Material: Silicone coating applied on top of the foam.

  - Function: Ensures water resistance, allowing use during aquatic activities such as swimming or bathing in a hot tub.

  - Protects the foam core from wear and tear caused by moisture.

- Front Velcro Patch:

  - Material: Sewn into the silicone layer for durability.

  - Function:

    - Enables the collar to roll up and pack compactly for easy storage.

    - Provides secure fastening when the collar is folded.

- Back Velcro Adjustment Strap:

  - Material: Sewn into the silicone and reinforced for frequent adjustments.

  - Function:

    - Allows the user to customize the collar size to fit snugly around the neck.

    - Provides a reliable fastening mechanism that resists slippage.

- Two-Part Epoxy Adhesive:

  - Usage: Bonds the polyurethane foam pieces for a strong and durable collar.

  - Function:

    - Ensures a seamless and durable connection between the individual parts

    - Prevents detachment even under extended use and environmental exposure.

## Exiting the System

- Rolling:

  - Lay the collar flat on a clean surface with the inner side facing upward.

  - Begin rolling the collar from one end toward the other, ensuring a tight and compact shape.

- Securing:

  - Use the front Velcro patch to fasten the rolled-up collar securely.

  - Ensure the Velcro is properly aligned to prevent the collar from unrolling during storage.

# Using the System

The following sub-sections provide detailed, step-by-step instructions on how to use the various functions or features of the Waterproof Soft Cervical Collar.

## Function: Adjustable Support for Neck Stabilization

The waterproof soft cervical collar provides adjustable support for stabilizing the neck during recovery, posture correction, or physical activities. The system incorporates user-friendly features to ensure comfort, water resistance, and versatility.

Figure 3. Soft cervical collar - Back view

## Function: Breathability Enhancement via Front Hole Design

- Purpose: Improves airflow to reduce discomfort during extended use.

- Expected Behavior:

  - Ensures the user feels cool and comfortable, especially in warm conditions.

  - Minimizes sweat accumulation around the neck.

- Special Instructions:

  - Align the breathable hole with the front of the neck for maximum ventilation.

- Caveats/Exceptions:

  - Misalignment during setup may reduce its effectiveness.

Figure 4. Soft cervical collar - Front view (red arrows indicate airflow)

## Function: Rolling Up the Waterproof Soft Cervical Collar

The roll-up feature is designed to enhance the collar's portability and storage convenience. By enabling the collar to compactly roll up, users can easily transport and store it without compromising its shape or functionality.

Figure 5. Soft cervical collar - Dimensions

# Troubleshooting & Support

The following sub-sections will describe the instructions to be taken when troubleshooting the soft cervical collar and provides information on maintenance and support.

## Error Conditions and Recovery Procedures

### Loose Fit Around the Neck

- Error Condition: The collar feels loose and does not provide adequate support.

- Likely Cause:

  - Velcro straps are not secured tightly.

  - Velcro fastener is worn or has lost its grip.

- Corrective Actions:

  - Adjust and secure the Velcro straps, ensuring a snug fit.

  - If the Velcro has weakened, clean it using a soft brush to remove dirt or debris.

  - If the issue persists, consider replacing the Velcro fasteners with new ones (contacting someone who specializes in sewing is recommended)

### Collar Does Not Stay Rolled Up

- Error Condition: The collar unrolls after being secured.

- Likely Cause:

  - Misalignment of the Velcro patch during rolling.

  - Worn-out Velcro or improperly applied silicone layer on Velcro.

- Corrective Actions:

  - Re-roll the collar carefully, ensuring the Velcro aligns correctly.

  - Clean the Velcro patch to restore grip.

  - Inspect the silicone coating; if it covers the Velcro, carefully remove excess material

### Tears or Cracks in the Silicone Coating

- Error Condition: Cracks or tears appear on the waterproof silicone layer.

- Likely Cause:

  - Excessive stress on the collar during use or storage.

  - Long-term wear and environmental factors

- Corrective Actions:

  - Apply a silicone repair adhesive (available at hardware or online stores) to the affected area.

  - Allow the adhesive to cure as per the manufacturer’s instructions.

  - Avoid using the collar until the repair is complete to prevent further damage.

### Foam Core Deformation

- Error Condition: The collar no longer holds its shape or becomes uneven.

- Likely Cause:

  - Excessive compression during storage.

  - Long-term exposure to moisture.

- Corrective Actions:

  - Take a large syringe and fill it up with air

  - Make a small incision in the silicone, refer to 5.1.1.3 on how to patch this incision after

  - Inject the cervical collar with air until the foam regains its structure

### Velcro Detachment

- Error Condition: Velcro fasteners detach from the silicone layer.

- Likely Cause:

  - Threads have detached from the Velcro patch and silicone layer.

- Corrective Actions:

  - Take a thread a needle

  - Carefully reinsert the detached Velcro by sewing it into the silicone

## Special Considerations

### For Water Use

- Rinse the collar with fresh water immediately after exposure to chlorinated or salt water.

- Dry thoroughly to avoid mold or material degradation.

### For Long-Term Storage

- Avoid rolling the collar tightly for extended periods to preserve the foam core’s elasticity.

- Store in a dry, cool place away from direct sunlight or heat.

## Maintenance for Error Prevention

- Regularly clean the Velcro fasteners to prevent dirt accumulation.

- Inspect the silicone coating and foam for signs of wear or damage.

- Avoid over-tightening or overstretching the collar during use or storage.

These recovery and error correction procedures ensure the collar remains functional and effective, even after prolonged use or minor issues.

## Emergency Contact

If there are any emergencies, please contact:

- Rayhan Karim: akari033@uottawa.ca

- Gavin Garfinkel: ggarf079@uottawa.ca

- Aisven Devanand: adeva042@uottawa.ca

- Matthew George: mgeor060@uottawa.ca

- Jonah Philip: jphil@uottawa.ca

# Product Documentation

### BOM (Bill of Materials)

Table 3. Bill of Materials

| Item Name | Description | Units of Measure | Quantity | Unit Cost | Extended Cost | Link |
| --- | --- | --- | --- | --- | --- | --- |
| Polyurethane foam | 16x12 inch sheet, 1.5 inches thick | Sheet | 2 | $11.50 | $23 | [1] |
| Waterproof Heavy-duty strips | Adhesive Velcro strips to attach anywhere, 1 inch x 24 feet | Strip | 1 | $12 | $12 | [2] |
| Silicone Rubber | Liquid silicone kit to pour over the polyurethane foam, 21.16 Oz two-part bottles | Kit | 1 | $23 (with prime) - $26 without | $23 (with prime) - $26 without | [3] |
| Total Product Cost (without taxes or shipping + prime membership) |  |  |  |  | $58 |  |
| Total Product Cost (with taxes + shipping + prime membership) |  |  |  |  | $65.54 |  |

### Equipment list

The following sub-sections separate the equipment list into six categories, these are:

- Cutting and Shaping Tools

- Adhesive Application Tools

- Sewing Tools

- Measurement and Alignment Tools

- Finishing and Testing Tools

- General Supplies

#### Cutting and Shaping Tools

- Utility Knife: For cutting the polyurethane foam into the desired shape.

- Scissors: For trimming silicone sheets and Velcro fasteners.

- Rotary Cutter: For precise cuts in the foam or silicone layer, if required.

#### Adhesive Application Tools

- Two-Part Epoxy Adhesive Kit: For bonding the silicone layer to the foam.

- Small Mixing Tray: To prepare and mix the epoxy adhesive and silicone

- Disposable Brushes or Applicators: For spreading the adhesive evenly.

#### Sewing Tools

- Heavy-Duty Sewing Machine: For attaching Velcro fasteners to the silicone layer.

- Thread (Nylon or Polyester): Strong, waterproof thread to withstand repeated adjustments.

- Needles: Heavy-duty needles compatible with silicone and Velcro materials.

#### Measurement and Alignment Tools

- Ruler or Measuring Tape: For measuring the foam, silicone layer, and Velcro dimensions accurately.

- Marking Pen or Chalk: To outline cutting and sewing points on materials.

#### Finishing and Testing Tools

- Heat Gun: To seal edges of silicone or reinforce adhesive bonds (optional).

- Clamps or Weights: For holding glued parts in place while the adhesive sets.

- Water Basin: For testing the waterproofing capabilities of the collar.

#### General Supplies

- Cleaning Materials:

  - Isopropyl alcohol or cleaning solution for preparing surfaces before adhesion.

  - Soft cloths for cleaning and drying the finished collar.

- Protective Gear:

  - Gloves and masks to handle adhesive safely and avoid skin contact.

### Instructions

The following sub-sections will cover the detailed step-by-step process done to make the final prototype of the soft cervical collar.

#### Step 1: Prepare the Materials

- Gather the following components:

  - Polyurethane foam sheet.

  - Silicone that comes in Part A and B for waterproofing.

  - Velcro strips (front patch and back fastener).

  - Two-part epoxy adhesive.

- Clean all materials using isopropyl alcohol to remove dust or grease.

#### Step 2: Cut the Polyurethane Foam

- Measure and Mark:

  - Using a measuring tape, mark the dimensions for the collar based on the user’s neck size and shape.

  - Typically, the length should wrap around the neck, and the height should provide adequate support without obstructing movement.

- Cut the Foam:

  - Use a utility knife or rotary cutter to carefully cut the foam along the marked lines.

  - Smooth any rough edges with scissors or a foam cutter.

#### Step 3: Apply the Silicone Coating

- Mix Silicone Parts:

  - Mix Parts A and B equally by weight or volume, in a mixing tray.

- Pour the Silicone:

  - Carefully place the mixed silicone onto the foam, very small amounts at a time.

  - Use a disposable brush to apply a thin, even layer of silicone to one side of the foam.

  - Repeat for the other side of the foam.

#### Step 4: Attach the Velcro

- Prepare Velcro Strips:

  - Cut Velcro into appropriate lengths: one patch for the front (storage) and one for the back (size adjustment).

- Sew Back Velcro:

  - Use a heavy-duty sewing machine to stitch the adjustable Velcro strip onto the silicone layer at the back of the collar. Ensure the stitching is tight and secure.

- Sew Front Velcro:

  - Use a heavy-duty sewing machine to stitch the adjustable Velcro strip onto the silicone layer at the front of the collar. Ensure the stitching is tight and secure.

#### Step 5: Final Assembly and Testing

- Inspect the Collar:

  - Check for loose edges, adhesive gaps, or misaligned components. Reinforce as needed.

- Test Waterproofing:

  - Submerge the collar in a water basin to verify the silicone layer’s effectiveness.

  - Ensure no water seeps into the foam core.

- Test Adjustability:

  - Wear the collar to ensure the Velcro fasteners function correctly for size adjustment and rolling.

#### Step 6: Finishing Touches

- Trim any excess material from the silicone or foam for a clean finish.

- Clean the entire collar using a damp cloth to remove adhesive residue.

## Testing & Validation

Many tests were done on all of the prototypes to see if the target specifications would be met by the final prototype. The tables below summarize all of the tests done for each prototype. All three prototypes can be seen in Figure 6.

Figure 6. Prototypes developed for soft cervical collar (1: Left - 2: Center - 2.5: Right)

Table 4. Summary of tests for Prototype 1

| Test Number | Test Name | Method of Data Collection | Test Plan | Test Results |
| --- | --- | --- | --- | --- |
| 1.1 | Rigidity and Flexibility | Qualitative | The prototype is compressed and folded when: i) without breathability holes, ii) with breathability holes and iii) with Vinyl sheet layer to see its rigidity and flexibility. | Flexible structure. Reduced rigidity with 3 holes added to the front. Reduced flexibility with Vinyl sheet layer added. |
| 1.2 | Compressibility | Quantitative | The height of the prototype laid flat will be measured before and after placing a 10 lbs weight on top and the difference in height will be recorded. | Prototype with initial height of 40 mm compressed by 20 mm with the 10 lbs weight on top. |
| 1.3 | Velcro Strength Performance | Qualitative/ Quantitative | Repeated removal and reconnection of the Velcro’s hoops and loops to see if it will deform under water. The collar will also be submerged in water for 10 minutes and the Velcro will be retested. | The Velcro kept the same strength and showed no signs of wear or tear throughout the repeated removal of the strips. Water made no difference to the Velcro’s strength. |
| 1.4 | Rollability & Portability | Quantitative | The prototype will be rolled up and compressed five times. The dimensions will be recorded. | When rolled up, the prototype is 6 inches in diameter and 4 inches in height. Holds shape due to flexibility. |
| 1.5 | Performance in Water | Qualitative/ Quantitative | The prototype will be submerged in water for 45 minutes and will then be tested for water retention and weight differences. | The vinyl sheet was not fully waterproof and allowed water to seep through – could not perform other water tests (Chlorine Water, Washing Machine). |

Table 5. Summary of tests for Prototype 2

| Test Number | Test Name | Method of Data Collection | Test Plan | Test Results |
| --- | --- | --- | --- | --- |
| 2.1 | Rigidity and Flexibility | Qualitative | The prototype is compressed and folded to see its rigidity and flexibility. | Very rigid structure and not flexible. |
| 2.2 | Compressibility | Quantitative | The height of the prototype laid flat will be measured before and after placing a 10 lbs weight on top and the difference in height will be recorded. | Prototype with initial height of 40 mm compressed by 0 mm with the 10 lbs weight on top. |
| 2.3 | Velcro Adhesive Performance | Qualitative | The adhesive side of the Velcro will be sticked on the Silicone surface and pulled on using the other side of the Velcro to try and detach it. | Adhesive did not stick to Silicone Rubber surface. Easily detached when lightly pulled. |
| 2.4 | Rollability & Portability | Quantitative | The prototype will be rolled up and compressed five times. The dimensions will be recorded. | Could not test as the structure was not flexible and also due to prototype not being collar sized. |
| 2.5 | Performance in Water | Qualitative/ Quantitative | The prototype is submerged in water for 45 minutes and it is then tested for water retention and weight differences. | Zero structural deformities and exact same weight after being submerged in water – Did not test in Chlorine Water or Washing Machine since prototype structure did not follow initial concept for final design due to failure in manufacturing process. |

Table 6. Summary of tests for Prototype 2.5

| Test Number | Test Name | Method of Data Collection | Test Plan | Test Results |
| --- | --- | --- | --- | --- |
| 3.1 | Rigidity and Flexibility | Qualitative | The prototype is compressed and folded to see its rigidity and flexibility. | Stiffer than prototype 1 while more flexible than prototype 2. Holds its shape after being compressed. |
| 3.2 | Compressibility | Quantitative | The height of the prototype laid flat will be measured before and after placing a 10 lbs weight on top and the difference in height will be recorded. | Prototype with initial height of 40 mm compressed by 5 mm with the 10 lbs weight on top. |
| 3.3 | Velcro Adhesive and Nylon Thread Performance | Qualitative | The adhesive side of the Velcro will be stuck on the Silicone surface and secured using Nylon threads sewed through the prototype. The Velcro will then be pulled on using the other side of the Velcro to try and detach it. | Velcro stayed secured to the Silicone Rubber surface using Nylon threads, however threads failed when high loads were applied on Velcro strips when pulling. Strategy for final prototype will be to use Nylon threads and coat edges of Velcro with Silicone Rubber and/or Glue. |
| 3.4 | Rollability & Portability | Quantitative | The prototype will be rolled up and compressed five times. The dimensions will be recorded. | Could not test as the prototype was not collar sized. |
| 3.5 | Performance in Water | Qualitative/ Quantitative | The prototype will be submerged in water for 45 minutes and will then be tested for water retention and weight differences. | Zero structural deformities and exact same weight after being submerged in water. |
| 3.6 | Performance in Chlorinated Water | Qualitative/ Quantitative | 0.05 mL of bleach will be mixed into 1 L of water to achieve concentration of chlorinated water. The prototype will be submerged in solution for 45 minutes and will then be tested for liquid retention & weight differences and inspected for surface deformities. | No visible damaged to Silicone surface and liquid did not seep inside prototype. |
| 3.7 | Performance inside Washing Machine | Qualitative/ Quantitative | The prototype will be placed in the washing machine and run for 5 machine cycles at 40°C. The prototype will then be inspected for surface deformities and tested for liquid retention & weight differences | No visible damaged to Silicone surface and liquid did not seep inside prototype. |

The compressibility for tested for all three prototypes as well as the Polyurethane foam by itself was summarized in Figure 7.

Figure 7. Compressibility of collar structures

After testing and validating the results, it was seen that prototypes 1 and 2.5 passed most of their tests and would form the basis of the final prototype. The development of the final prototype combined the general shape of prototype 1 with the structure of prototype 2.5.

# Conclusions and Recommendations for Future Work

In conclusion, this project served to expand the untapped market of the soft cervical collar used by individuals with neck mobility impairments during physical activities as it provided various challenges to overcome. The final prototype that was developed combined several engineering concepts and techniques to solve this important problem. Other avenues that could potentially benefit similar attention could be joint supporting collars that can be used under water or in environments demanding physical activity.

Given more time and money, the Alien Collars Co team was planning to develop a cloth layer made of sportswear or wetsuit fabric that could be detachable with the help of Velcro strips placed in various spots across the collar. This layer would provide a better skin comfort and reduce friction from the Silicone surface. A preliminary conceptual drawing of this idea can be seen in Figure 8. Appropriate foam cutting tools would have also been utilized in order to achieve a final product that’s aesthetically pleasing, and more research would be done on the Silicone to get perfect the curing process and find improvements in comfort.

Figure 8. Fabric layer conceptual drawing

# Bibliography

| [1] | “2 PCs Polyurethane Foam Sheet- 16 x 12 x 1.5 Inch Cuttable Foam Inserts for Cases- Packing Foam Pads for Toolbox Camera Storage and Crafts : Amazon.Ca: Tools & home improvement,” Amazon.ca. [Online]. Available: Link. |
| --- | --- |
| [2] | “2 Roll Hook and Loop Tape, Heavy Duty Dual Lock, Nylon Self Adhesive Dual Lock, 1 Inch x 26 Feet Reusable Hook and Loop Interlocking Mounting Tape Strips Fastener for Home Office Carpet Mats : Amazon.Ca: Home,” Amazon.ca. [Online]. Available: Link. |
| [3] | “BBDINO Silicone Mold Making Kit, Super Elastic Mold Making Silicone Rubber, Liquid Silicone for Mold Making, N.w. 21.16 oz, 1:1 by Volume, Ideal for Casting 3D Silicone Molds,” Amazon.ca. [Online]. Available: Link. |

APPENDICES

# APPENDIX I: Design Files

The files related to this manual can be found in Table 7.

Table 7. Referenced Documents

| Document Name | Document Location and/or URL | Issuance Date |
| --- | --- | --- |
| GNG2101 Group B2.3 - Team_Contract - PD A.pdf | https://makerepo.com/akari033/2182.soft-cervical-collar-by-alien-collars-co | 2024/12/02 |
| PD_B-D_Group_B2.3.docx | https://makerepo.com/akari033/2182.soft-cervical-collar-by-alien-collars-co | 2024/12/02 |
| PD_E-I_Group_B2.3.docx | https://makerepo.com/akari033/2182.soft-cervical-collar-by-alien-collars-co | 2024/12/02 |
| Waterproof neck collar - Concept 1.SLDPRT | https://makerepo.com/akari033/2182.soft-cervical-collar-by-alien-collars-co | 2024/12/02 |
| Waterproof neck collar - Detailed Design 2.SLDPRT | https://makerepo.com/akari033/2182.soft-cervical-collar-by-alien-collars-co | 2024/12/02 |
| GNG2101 B2 3 PD I Video | https://youtube.com/shorts/1GSmvYLmCD0 | 2024/12/02 |